Response to Reviewers:

Reviewer C:

We appreciate the reviewer for the positive review and for providing valuable comments for further improvement.

Major correction.

In Line – 145. “The ratio between MLC and jaw field measurements had an average range across FSs of 0.019 for FFout beams and 0.045 for FFin beams, with a peak value at 4 x 4 cm2” The values of 0.019 and 0.045 should be re-checked which does not comply with the graphical representation in Fig-4.

Response:

The values reported in Line 145 were calculated by averaging, across energy and SSD, the ranges in MLC-to-Jaw measurements as FS was varied from 2 cm to 8 cm square field length. The values do not comply with Fig-4 because only 100 cm SSD is shown in that figure. Line 145 was revised for clarity to:

"At all SSDs, the ratio between MLC and jaw field measurements was relatively stable

across FSs for FFout beams (average arithmetic range in ratios across FS of 0.019,

taken across all SSDs and energies); whereas much larger variability was seen for FFin

beams (average arithmetic range in ratios across FS of 0.045)."

Minor correction.

Line – 130. “For example, at 2 x 2 cm2 jaw FS, surface dose ratios differed from 80 cm SSD to 100 cm SSD by 0.8% and 0.0% for 6 MV and 10 MV, respectively. However, at 20 x 20 cm2 FS, these differences in ratios increased to 4.5% and 6.9%. Similar trends were observed for MLC fields.” It would be better to enlarge the scale in a separate figure to show the differences for small and large field sizes and SSD for better clarity.

Response:

Figure 5 will be asked to be printed in enlarged size fitting into one full page.

In section III.D. Reference Depth Dose Comparison it would be better to mention that in particular for Truebeam the FFout beams were not tuned to FFin beams at d10 like other vendor Linac.

Response:

The following sentence was added. “Because no beam tuning to match FFout beams with their corresponding FFin beams was made for the TrueBeam, the dose difference at d10 is expected even for the 10 x 10 cm2 reference field.”

Reviewer D:

We appreciate the reviewer for the review and for providing valuable comments for improvement.

The authors are presenting a measurement series of skin dose with variable SSD, Energy, Field Sizes and FFF / FF Beam and find that FFF delivery has a higher relative skin dose compared to FF delivery, increasing with higher energy. While the measurements are all fine and presented in a straight forward manner, the work lacks of clinical relevance despite the promise in the title and unfortunately novelty as well.

1) The measurements presented in this work can also be subtracted from comprehensive commissioning data for FFF Linacs and the results become immediately obvious, hence the findings are not new.

Response:

We agree with the reviewer that, in principle, it may be possible to extract skin doses accurately from commissioning data. However, depending on the situation, it may not be the case in reality. The most common method of skin dose measurement in actual practice is to use a parallel-plate chamber with the necessary correction applied, which is the next best choice after expensive and exhausting extrapolation chamber measurement. However, to our best knowledge, many users utilize a cylindrical chamber for beam scanning during their commissioning. And it is well known that cylindrical chambers suffer from significant volume averaging effects in high-gradient regions (e.g. build-up regions) which is worst at the surface; thus, skin dose estimation from cylindrical chamber data is not reliable. In addition, even in the case that someone can potentially extract skin doses from the commissioning data for his/her clinic, we believe it is still valuable to have such data publicly available. This issue has been addressed in the discussion of the revision.

2) Why would higher relative skin dose be clinically relevant other than for the fact that FFF modes should be used only in combination with more fields or arcs as compared to standard FF modes as to not violate skin dose constraints (which are well known, too)? This fact seem obvious and generally FFF modes are only used in combination with SBRT treatment, where due to high single doses multiple fields or arcs are used anyway (i.e., to generate a steeper drop off in all directions).

Response:

We agree with the reviewer that greater benefit of using FFF mode can be obtained from SBRT applications as SBRT commonly utilizes either multiple or arc beams. However, that does not necessarily guarantee that there will be no skin complication. In fact, skin complication can be severe in SBRT because of its nature of hypo-fractionated high-fractional dose, and severe skin complications have been reported with SBRT. In addition, from time to time, the desire to have a greater number of beams is compromised with the necessity of beam path avoidance for OARs, resulting in increased skin dose in SBRT. This issue has been addressed in the discussion of the revision.

Comment: The sentences “where dose is almost exclusively delivered by small fields using [] (MLCs) …” and Early investigations into the dosimetric properties …” are surely incorrect as to the fact that most SBRT or SRS is performed on dedicated machines such as Gamma Knife and CyberKnife and more importantly the Gamma Knife and the CyberKnife use FFF beams (i.e., for CyberKnife it is 6 MV) since 1968 and 1994, respectively. Hence, FFF beams are not new and the dosimetry of said beams is well characterized and published in countless papers.

Response:

The relevant statement has been omitted in the revision.