Reviewer B:

We appreciate the reviewer for the positive review and for providing valuable comments for further improvement.

The present work is an evaluation of skin doses from FFF treatments. Skin dose is of interest for FFF beams because it is one of the biggest concerns for this treatment modality. While there have been previous papers on this subject, many of the previous skin dose studies have been sub-optimal. Consequently, I believe there is room for important additional studies on skin dose from these beams.

Major comments

The authors should refer to the AAPM report on FFF beams and references therein:  Xiao et al, JACMP, 16(3):12-29 2015

Response:

This was added (explicitly noted as ‘Xiao’ in this review version at line 145 to indicate where it was referenced).

The authors should use nomenclature consistent with the AAPM report. That is, flattening filter free beams should be referred to as FFF. Traditional flattened beams can be referred to as “flattened” or something else at the authors’ discretion.

Response:

Changes in beam names were made as suggested. The figures were also updated accordingly.

The dose rate study is not correct. The TrueBeam changes rep rates by pulse dropping. That is, for a 6 FFF beam, 1400 MU/min and 600 MU/min have the same instantaneous dose rate (dose per pulse) and therefore the same dose rate effects. This study design doesn’t test anything. Either remove this section, or if the dose rate effects of this dosimeter warrant characterization (quite possible), then this should be done following recommended evaluation of recombination per the FFF report or the AAPM TG-51 Addendum (McEwan et al) and if necessary, by varying the dose rate (e.g., by changing the instantaneous dose rate via SSD/depth variation).

Response:

The dose rate portion of this study was removed (previously section III A).

Uncertainty analysis should be included in this study. What are the error bars associated with the data in each of the figures?

Response:

An uncertainty analysis section was added at the end of Results to discuss the measurement uncertainty. Based on the analysis, the uncertainty turned out insignificant (less than 0.2%).  For instance, the 2x2 cm2 field at 100 cm SSD had a FFF-to-flattened ratio of 1.4525. The uncertainty for this ratio (which was the largest of any field size) was found via error propagation to be 0.0016. Therefore, the error bars would appear as a single horizontal dash at each data point, and we decided not to include the error bars in the figures to avoid unnecessary business.

Minor comments

FS is not a standard acronym and should not be used. Please spell out “field size”.

Response:

“FS” was replaced with “field size” throughout the manuscript.

Abstract. The results section is not particularly clear.

Response:

The results section of the abstract was changed for clarity.

1)       “The ratio between MLC and jaw measurements had an average range across FS of….” This sentence is very hard to understand – I suggest you add details.

Response:

Results were changed for clarity.

2)      You only discuss results for FS <= 3x3. This is a very small subset of your FS results.

Response:

Results were changed for clarity, removing the sentence which singled-out < 3x3 cm2 results.

3)      Comments are made about dose rate, but no methodology is described. I suggest you remove this statement from the results unless you really want to focus on them.

Response:

Dose rate investigation was removed from this paper and is no longer mentioned in the abstract.

4)      You don’t discuss SSD variations. Move the conclusion statement to the results.

Response:

Your observation is valid. SSD variations are now mentioned. Also, abstract sections were removed.

Abstract. The conclusion section is a results/summary section. The abstract would be more clear if the results were in the results section, and the conclusion section presented the true conclusion of this study.

Response:

The abstract was changed to correct for this valid criticism. Sections were removed.

Introduction: “enabling beneficial techniques such as breath hold or active breathing control to be used for otherwise ineligible patients” This statement should have a reference.

Response:

Clarification provided. This change in eligibility is specific to our clinic; whether it also applies to other institutions will vary.

Ln 156 should have a reference

Response:

Clarification was provided. The authors meant to say that this was “discussed previously in the introduction of this paper” rather than “discussed previously in the literature”

Ln 163-164 please provide a reference – I have not seen this claim made previously. Also, I’m a little confused about “saturation of the monitor chamber” (which I don’t believe happened) being related to the number of electrons reaching the patient.

Response:

The sentence was slightly reworded, and a reference was provided.

Ln 184 – why was this peak only seen for the flattened 10 MV beam? If there is a balancing between these two effects, why wasn’t it seen with other modalities?

Response:

This paragraph was reworded for clarity, and a discussion was added in the Discussion section (Lines 290-298).